## Input Buffer
Arlyn Scales - EtherNet/IP: T2O INPUT BUFFER - LITTLE ENDIAN (LE) [CIP standard, DCBA byte order] (v2.2)
Byte offsets are within the T2O assembly payload (288 bytes / 0x120 total). Library handles CIP framing.
KIND | OFFSET (Hex) | OFFSET (Dec) | ITEM | TYPE | DESCRIPTION | NOTES
TOP PLATFORM (Base = 0x00)
Input Buffer 0x0000 to 0x001F "On Top" Platform T2O (Target to Originator) LITTLE ENDIAN (LSB -> MSB) | 0x0000 | 0 | Unit (Top Platform) | UINT8 | Unit code | 0=lb 1=kg 2=g 3=oz ...
| 0x0001 | 1 | [RESERVED] | UINT8 | Always 0
| 0x0002-0x0003 | 2-3 | Gross Weight [low WORD] | UINT16 LE | Low 16 bits of IEEE-754 float | DCBA: low word at lower address
| 0x0004-0x0005 | 4-5 | Gross Weight [high WORD] | UINT16 LE | High 16 bits of IEEE-754 float
| 0x0006-0x0007 | 6-7 | Net Weight [low WORD] | UINT16 LE
| 0x0008-0x0009 | 8-9 | Net Weight [high WORD] | UINT16 LE
| 0x000A-0x000B | 10-11 | Tare Weight [low WORD] | UINT16 LE
| 0x000C-0x000D | 12-13 | Tare Weight [high WORD] | UINT16 LE
| 0x000E-0x000F | 14-15 | Status 1 (Top Platform) | UINT16 LE | 16-bit weight status | See Weight Status sheet
| 0x0010-0x0011 | 16-17 | Displayed Value [low WORD] | UINT16 LE
| 0x0012-0x0013 | 18-19 | Displayed Value [high WORD] | UINT16 LE
| 0x0014-0x0015 | 20-21 | Status 2 (GLOBAL) | UINT16 LE | Setpoints + HI/OK/LO | See Weight Status sheet
| 0x0016-0x001C | 22-28 | [RESERVED] | - | Always 0 | NEW v2.2: extension fields moved to +0x1D-0x1F
| 0x001D | 29 | Parts Counting Active (On Top) | UINT8 | 0=Off 1=On (active + sample loaded) | NEW v2.2: moved; now 1-byte UINT8
| 0x001E | 30 | Front Panel Screen Type | UINT8 | 0=Single 1=Scoreboard 2=Multi-Platform 3=Diff/Total | NEW v2.2: moved; now 1-byte UINT8
| 0x001F | 31 | Current Platform Number (On Top) | UINT8 | 0-5 -> Platform 1-6 | NEW v2.2: moved; now 1-byte UINT8
TOTALS (Base = 0x20)
[MULTI-PLATFORM SCALES ONLY] Input Buffer 0x0020 to 0x003F (Target to Originator) LITTLE ENDIAN | 0x0020 | 32 | Unit (Totals) | UINT8 | Same encoding as Top Platform
| 0x0021 | 33 | [RESERVED] | UINT8 | Always 0
| 0x0022-0x0023 | 34-35 | Total Gross Weight [low WORD] | UINT16 LE
| 0x0024-0x0025 | 36-37 | Total Gross Weight [high WORD] | UINT16 LE
| 0x0026-0x0027 | 38-39 | Total Net Weight [low WORD] | UINT16 LE
| 0x0028-0x0029 | 40-41 | Total Net Weight [high WORD] | UINT16 LE
| 0x002A-0x002B | 42-43 | Total Tare Weight [low WORD] | UINT16 LE
| 0x002C-0x002D | 44-45 | Total Tare Weight [high WORD] | UINT16 LE
| 0x002E-0x002F | 46-47 | Status 1 (Totals) | UINT16 LE | Bit 0=stable when ALL platforms stable | Other bits unused
| 0x0030-0x0035 | 48-53 | [RESERVED] | - | Always 0
| 0x0036-0x003C | 54-60 | [RESERVED] | - | Always 0
| 0x003D | 61 | Parts Counting Active (Totals) | UINT8 | 0=Off 1=On (active + sample loaded) | NEW v2.2: moved; now 1-byte UINT8
| 0x003E | 62 | Front Panel Screen Type | UINT8 | 0=Single 1=Scoreboard 2=Multi-Platform 3=Diff/Total | NEW v2.2: moved; now 1-byte UINT8
| 0x003F | 63 | Current Platform Number (On Top) | UINT8 | 0-5 -> Platform 1-6 | NEW v2.2: moved; now 1-byte UINT8
PER-PLATFORM BLOCKS - Platform N at base (0x40 + N x 0x20) where N=0..5, [N=0 -> PLATFORM 1], [N=1 -> PLATFORM 2] etc.
P0=0x40 P1=0x60 P2=0x80 P3=0xA0 P4=0xC0 P5=0xE0
(Target to Originator) LITTLE ENDIAN | base+0x00 | - | Unit (Platform N) | UINT8
| base+0x01 | - | [RESERVED] | UINT8 | Always 0
| base+0x02-0x03 | - | Gross Weight [low WORD] | UINT16 LE | Low 16 bits of IEEE-754 float | DCBA: low word at lower address
| base+0x04-0x05 | - | Gross Weight [high WORD] | UINT16 LE | High 16 bits of IEEE-754 float
| base+0x06-0x07 | - | Net Weight [low WORD] | UINT16 LE
| base+0x08-0x09 | - | Net Weight [high WORD] | UINT16 LE
| base+0x0A-0x0B | - | Tare Weight [low WORD] | UINT16 LE
| base+0x0C-0x0D | - | Tare Weight [high WORD] | UINT16 LE
| base+0x0E-0x0F | - | Status 1 (Top Platform) | UINT16 LE | 16-bit weight status | See Weight Status sheet
| base+0x10-0x11 | - | Displayed Value [low WORD] | UINT16 LE
| base+0x12-0x13 | - | Displayed Value [high WORD] | UINT16 LE
| base+0x14-0x15 | - | Status 2 (GLOBAL) | UINT16 LE | Single-platform only; 0 in multi-platform mode | Read from Top Platform block (0x0014) in multi-platform mode
| base+0x16-0x1F | - | [RESERVED] | - | Always 0
PLATFORM 0 BACKWARD-COMPAT BLOCK - base 0x100
Duplicate of Platform 0 (base 0x40) for backward compatibility from previous firmware.
(Target to Originator) LITTLE ENDIAN | base+0x00 | - | Unit (Platform N) | UINT8
| base+0x01 | - | [RESERVED] | UINT8 | Always 0
| base+0x02-0x03 | - | Gross Weight [low WORD] | UINT16 LE | Low 16 bits of IEEE-754 float
| base+0x04-0x05 | - | Gross Weight [high WORD] | UINT16 LE | High 16 bits of IEEE-754 float
| base+0x06-0x07 | - | Net Weight [low WORD] | UINT16 LE
| base+0x08-0x09 | - | Net Weight [high WORD] | UINT16 LE
| base+0x0A-0x0B | - | Tare Weight [low WORD] | UINT16 LE
| base+0x0C-0x0D | - | Tare Weight [high WORD] | UINT16 LE
| base+0x0E-0x0F | - | Status 1 (Top Platform) | UINT16 LE | 16-bit weight status
| base+0x10-0x11 | - | Displayed Value [low WORD] | UINT16 LE
| base+0x12-0x13 | - | Displayed Value [high WORD] | UINT16 LE
| base+0x14-0x15 | - | Status 2 (GLOBAL) | UINT16 LE | Single-platform only; 0 in multi-platform mode | Read from Top Platform block (0x0014) in multi-platform mode
| base+0x16-0x1F | - | [RESERVED] | - | Always 0
Copyright © Arlyn Scales, 2026

## Output Buffer
Arlyn Scales - EtherNet/IP: O2T OUTPUT BUFFER - Originator -> Target (Commands)
One byte per command address. Rising-edge triggered: write non-zero to fire, then clear to 0x00. Run/Idle header (4 bytes) prepended by PLC is stripped by firmware - not counted in these offsets. O2T assembly size: 288 bytes (9 blocks x 32 bytes, mirroring T2O layout).
KIND | OFFSET (HEX) | OFFSET (DEC) | ITEM | ACTION | DESCRIPTION | NOTES
TOP PLATFORM
Output Buffer [0x0000 to 0x000F] O2T (Originator to Target) LITTLE ENDIAN Set "Bit 0" on Each Byte to Execute Function | 0x0000 | 0 | Zero | 1: Execute | Perform zero on this platform
| 0x0001 | 1 | [RESERVED] |  | Future use - always write 0
| 0x0002 | 2 | Tare | 1: Execute | Perform tare on this platform
| 0x0003 | 3 | Clear Tare | 1: Execute | Clear tare back to zero
| 0x0004 | 4 | Change Unit | 1: Toggle to next active unit | Cycle to next unit on this platform
| 0x0005 | 5 | Net/Gross Toggle | 1: Net 0: Gross | Toggle net <-> gross display on this platform | 1: Toggle Net/Gross
| 0x0006 | 6 | Cycle Setpoint | 1: Toggle Cycle (Abort Cycle) | Cycle setpoint (ignored in flow-rate mode) | 1: Toggle
| 0x0007 | 7 | Avg Weight & Hold | 1: Start 0: Stop | Toggle Average Weight & Hold mode | 1: Toggle
| 0x0008 | 8 | Flow Rate Cycle | 1: Start 0: Stop | Cycle setpoint in flow-rate mode only | 1: Toggle
| 0x0009 | 9 | Log Weight | 1: Execute 0: Stop | Log weight entry
| 0x000A | 10 | Print | 1: Execute 0: Stop | Print
| 0x000B | 11 | Select OnTop |  | Make this platform the active OnTop display
| 0x000C-0x001F | 12-31 | [RESERVED] |  | Future use - always write 0
MULTIPLATFORM
MULTIPLATFORM: PER-BLOCK COMMAND TEMPLATE (applies to each of the 9 blocks; see Block Address Map below)
Output Buffer [Base] + [0x0000 to 0x000F] O2T (Originator to Target) LITTLE ENDIAN Set "Bit 0" on Each Byte to Execute Function | base+0x00 | 0 | Zero |  | Perform zero on this platform | Reserved - not yet implemented
| base+0x01 | 1 | [RESERVED] |  | Future use - always write 0
| base+0x02 | 2 | Tare |  | Perform tare on this platform | Reserved - not yet implemented
| base+0x03 | 3 | Clear Tare |  | Clear tare back to zero | Reserved - not yet implemented
| base+0x04 | 4 | Change Unit |  | Cycle to next unit on this platform | Implemented
| base+0x05 | 5 | Net/Gross Toggle |  | Toggle net <-> gross display on this platform | Reserved - not yet implemented
| base+0x06 | 6 | Cycle Setpoint |  | Cycle setpoint (ignored in flow-rate mode) | GLOBAL - same effect regardless of which block it arrives in
| base+0x07 | 7 | Avg Weight & Hold |  | Toggle Average Weight & Hold mode | GLOBAL - same effect regardless of which block it arrives in
| base+0x08 | 8 | Flow Rate Cycle |  | Cycle setpoint in flow-rate mode only | GLOBAL - same effect regardless of which block it arrives in
| base+0x09 | 9 | Log Weight |  | Log weight entry | GLOBAL - same effect regardless of which block it arrives in
| base+0x0A | 10 | Print |  | Print | GLOBAL - same effect regardless of which block it arrives in
| base+0x0B | 11 | Select OnTop |  | Make this platform the active OnTop display | Reserved - not yet implemented
| base+0x0C-0x1F | 12-31 | [RESERVED] |  | Future use - always write 0
BLOCK ADDRESS MAP
BLOCK ADDRESS MAP (absolute base addresses in the 288-byte O2T assembly), P0 = Platform 1, P1 = Platform 2, etc.
| BLOCK | ASSEMBLY | ACTS ON |  | NOTES
| Block 0 (OnTop) | 0x0000 | Currently-displayed (OnTop) platform |  | Check T->O Block 0 field 0x001F for active platform number
| Block 1 (Totals) | 0x0020 | Currently-displayed (OnTop) platform |  | Per-platform commands act same as Block 0
| Block 2 (P0) | 0x0040 | Platform 0 |  | Reserved - not yet implemented
| Block 3 (P1) | 0x0060 | Platform 1 |  | Reserved - not yet implemented
| Block 4 (P2) | 0x0080 | Platform 2 |  | Reserved - not yet implemented
| Block 5 (P3) | 0x00A0 | Platform 3 |  | Reserved - not yet implemented
| Block 6 (P4) | 0x00C0 | Platform 4 |  | Reserved - not yet implemented
| Block 7 (P5) | 0x00E0 | Platform 5 |  | Reserved - not yet implemented
| Block 8 (P0 compat) | 0x0100 | Platform 0 |  | Backward-compat alias for Block 2 - targets P0 regardless of OnTop
NOTE: Blocks 0 and 1 target the currently-displayed (OnTop) platform for per-platform commands (Zero, Tare, Unit, etc.). Global commands (Cycle, Avg Hold, Flow Rate, Log, Print) produce identical effect from any block.
Copyright © Arlyn Scales, 2026
SPECIAL NOTES ON SCREENS |  |  | SINGLE PLATFORM SCREEN | All Coil commands will work "as is".
|  |  | SCOREBOARD SCREEN | Coil commands will work only if buttons are visible
|  |  | TOTAL SCREEN | All Coil commands will work "as is".

## Weight Status
Arlyn Scales - EtherNet/IP: WEIGHT STATUS REGISTER BIT MAPS
STATUS 1 - 16-bit UINT16 LE (Top Platform: 0x000E ¦ Per-Platform: base+0x0E ¦ Totals: 0x002E)
KIND | BIT | MASK (Hex) | DISCRETE REF. | ITEM | WHEN SET (=1) | NOTES
TOP PLATFORM
Status 1 - Top Platform Input Buffer Address: 0x000E | 0 | 0x0001 | 010001 | OK Indicator | Set when no motion AND no overload (cleared otherwise)
| 1 | 0x0002 | 010002 | Net Center of Zero | "zr" on, net mode
| 2 | 0x0004 | 010003 | Gross Center of Zero | "zr" on, gross mode
| 3 | 0x0008 | 010004 | Net Display | Indicator showing net
| 4 | 0x0010 | 010005 | Gross Display | Indicator showing gross
| 5 | 0x0020 | 010006 | [RESERVED] | Firmware always returns 0 | Was: Zero in Progress
| 6 | 0x0040 | 010007 | [RESERVED] | Firmware always returns 0 | Was: Tare in Progress
| 7 | 0x0080 | 010008 | [RESERVED] | Firmware always returns 0 | Was: Average & Hold in Progress
| 8 | 0x0100 | 010009 | [RESERVED] | Firmware always returns 0 | Was: Weight Hold
| 9 | 0x0200 | 010010 | [RESERVED] | Firmware always returns 0 | Was: Zero Tracking
| 10 | 0x0400 | 010011 | No Motion | 1=No motion 0=Motion
| 11 | 0x0800 | 010012 | Stable | 1=Stable 0=Unstable
| 12 | 0x1000 | 010013 | Overload | 1=Overload 0=Normal
| 13 | 0x2000 | 010014 | Weighing Failure | Weight cannot be shown | Read register bit; discrete input not populated in firmware
| 14 | 0x4000 | 010015 | [RESERVED] | Firmware always returns 0 | Was: Error
| 15 | 0x8000 | 010016 | (empty)
STATUS 1 FOR MULTPLATFORM TOTALS: Only Bit 0 (OK Indicator) used - 1 when ALL attached platforms are stable. Discrete 012001 mirrors this bit.
MULTIPLATFORM
MULTIPLATFORM: Per-platform discrete addresses: 010N01-010N16 for Platform N where N=1, 2, 3 … 6.
STATUS 2 - GLOBAL 16-bit UINT16 LE (Top Platform: 0x0014 ¦ 0 in per-platform blocks in multi-platform mode)
Decode HI/OK/LO: mask with 0x0700, shift right 8 -> bit0=HI, bit1=OK, bit2=LO. Only one of HI/OK/LO set at a time.
KIND | BIT | MASK (Hex) | DISCRETE REF. | ITEM | WHEN SET (=1) | NOTES
Status 2 - GLOBAL STATUS (All Platforms) Input Buffer Address: 0x0014 | 0 | 0x0001 | 010017 | Setpoint 1 | On
| 1 | 0x0002 | 010018 | Setpoint 2 | On
| 2 | 0x0004 | 010019 | Setpoint 3 | On
| 3 | 0x0008 | 010020 | Setpoint 4 | On
| 4 | 0x0010 | 010021 | Setpoint 5 | On
| 5 | 0x0020 | 010022 | Setpoint 6 | On
| 6 | 0x0040 | 010023 | Setpoint 7 | On
| 7 | 0x0080 | 010024 | Setpoint 8 | On
| 8 | 0x0100 | 010025 | HI Output | Comparator high
| 9 | 0x0200 | 010026 | OK Output | Comparator within band
| 10 | 0x0400 | 010027 | LO Output | Comparator low
| 11 | 0x0800 | 010028 | [RESERVED] | Always 0
| 12 | 0x1000 | 010029 | [RESERVED] | Firmware always returns 0 | Was: User Input 1
| 13 | 0x2000 | 010030 | [RESERVED] | Firmware always returns 0 | Was: User Input 2
| 14 | 0x4000 | 010031 | [RESERVED] | Firmware always returns 0 | Was: User Input 3
| 15 | 0x8000 | 010032 | (empty)
Copyright © Arlyn Scales, 2026

## Connection Config
Arlyn Scales - EtherNet/IP: CONNECTION CONFIGURATION
CIP Assembly Parameters
PARAMETER | VALUE | NOTES
Connection Class | 1 | Class 1 - Implicit I/O Messaging
T2O Connection Point (Input Assembly) | 102 | Device -> Controller (scale data out)
O2T Connection Point (Output Assembly) | 101 | Controller -> Device (commands in)
Config Connection Point | 103
T2O Data Size | 288 bytes | Full assembly payload (0x120)
O2T Data Size | 288 bytes | 16 command addresses (0x00-0x0F)
Run/Idle Header | O2T only | 4-byte CIP header prepended by PLC; stripped by firmware
T2O Trigger | Cyclic | Device updates on each cycle
Float Encoding - Little-Endian DCBA
PARAMETER | VALUE | NOTES
Byte Order | CIP Little-Endian (DCBA) | Low word at lower byte address; each 16-bit word in LE byte order
PLC Tag Type | REAL | Standard CIP REAL - no byte swapping needed in RSLogix / Studio 5000
Example: 100.0 lb | 0x42C80000 | Bytes in buffer: 00 00 C8 42 (+0=0x00, +1=0x00, +2=0xC8, +3=0x42)
Assembly Address Map
BLOCK | BASE OFFSET | SIZE / NOTES
Top Platform | 0x0000 | 32 bytes - Active (On-Top) platform data
Totals | 0x0020 | 32 bytes - Multi-platform combined totals
Platform 0 | 0x0040 | 32 bytes - Physical platform 1
Platform 1 | 0x0060 | 32 bytes - Physical platform 2
Platform 2 | 0x0080 | 32 bytes - Physical platform 3
Platform 3 | 0x00A0 | 32 bytes - Physical platform 4
Platform 4 | 0x00C0 | 32 bytes - Physical platform 5
Platform 5 | 0x00E0 | 32 bytes - Physical platform 6
Platform 0 (compat) | 0x0100 | 32 bytes - Backward-compat duplicate of Platform 0
TOTAL | - | 288 bytes (0x120)
Copyright © Arlyn Scales, 2026

## Future Expansion
Arlyn Scales - FUTURE EXPANSION - Not Yet Implemented
# | Feature
1 | Parts Counting detail - per-platform count, sample weight, and PC status beyond active flag
2 | Setpoint write - push setpoint values from PLC/controller to device
3 | Tare write - send instant tare value from controller to device
4 | Sample write - send parts-counting sample weight from controller to device
5 | Raw gram weights - per-platform weight in grams regardless of display unit